Please Note: These instructions are for all members of the PX5 family of instruments and the EP1. The PX5 firmware will update ALL members of the PX5 family of instruments, including the PX5, PX5-400, PowerGuide/PG4400, and the PowerVisa.  The EP1 firmware will update ONLY the EP1

- Download the Firmware (Hostcode.bin) and the DFormat from the link you just received. This is the Dranetz FTP Site containing the firmware and the DFormat update utility.

- Place both/all files into the same folder.

- Connect a Compact Flash Card reader to your computer. Note what drive your computer assigns to the CF Card Reader.

- Insert the CD Card from your instrument into the CF Card reader.

- Double click on the DFormat utility.

- The DFormat utility should start and find your CF Card in the CF Card reader.  If the Utility does not find the CF Card in your reader please see the note below.

- Once the DFormat utility starts follow all instructions on the Computer for the DFormat utility.

- Once the DFormat utility has formatted and prepared the CF Card place the card back into the instrument and power cycle/power on the instrument.

- Follow all instructions on the instrument, including writing the firmware to memory.

- Once the Firmware has been installed into the instrument, remove the CF Card and restart the instrument without the CF Card installed.

- Confirm that the current firmware is in the instrument.  Firmware version is displayed top center of the Instrument home screen.

- Re-insert the CF Card and format the CF Card.

Note: As a safety precaution, the DFormat will only read a CF Card that is labeled /named DBMI32.  If the CF Card is not named DBMI32 then use Windows Explorer to rename the CF Card